## Slide 1
SPICA-VIS Design Review, July 23-24, 2020
Global view
Denis Mourard/Stéphane Lagarde
General Requirements
VIS
STS Six Telescope Simulator
Non resolved sources for alignment cophasing and calibration
LDC-COM Longitudinal Dispersion Compensator
Chromatic OPD compensation between IR and Visible
CHARA
IR
SPICA-FT SPICA Fringe Tracker
Fringe tracking
LDC-VIS Longitudinal Dispersion Compensator
Chromatic OPD compensation for visible light
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
SFU-INM Fiber Injection Optics
Optical system for the focalization of the beam at the fiber head level
SFU-OPF Optical Fiber
Optical Fiber for spatial filtering
SFU-RCM Fiber Reception Optics
Micro-Lens array for the beam collimation
Vgroove for beam set-up (3B–2B–4B–7B– B )
SCU Spectral Calibration Unit
Spectral sources for spectral Calibration
FBI Fiber Back Illumination
Light source to back illuminate to determine the reference axis of the fiber (on BCS)
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
SPICA-VIS
Alignment and Calibration Units
Beam Optimization Units
Spatial Filter Units
Spectro-Interferometer Units
Detector
Units external to SPICA-VIS

## Slide 2
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Feeding Optics (SFO-FPO)
Removable mirrors (motorized)
Spec. Repeatability: <10as & <1mm

## Slide 3
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Low frequency beam adjustment (SPO-IMG & BAS-PUP)
Slow tip/tilt mirrors (motorized)
Spec. Accuracy: <10as & <1mm

## Slide 4
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Fast frequency beam adjustment (SPO-IMG & BAS-PUP)
Fast tip/tilt mirrors (motorized)
Spec. Accuracy: <1as
Frequency: ≈ 200 Hz

## Slide 5
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Beam Control System (BCS)
Visible camera for beam control
Possibility to switch between image and pupil plans

## Slide 6
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Atmospheric Dispersion Correction (ADC)
Prisms (2 per beam, motorized)
Spec. Frequency: Every 10mns
Accuracy: 0.5°

## Slide 7
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Polarization Delay Compensation (PDC)
Birefringent plates (motorized)
Spec. Accuracy: 0.1°

## Slide 8
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Shutters (SHU)
Slow shutters (motorized)

## Slide 9
SPICA-VIS Design Review, July 23-24, 2020
Modules before the optical fibers
Denis Mourard/Stéphane Lagarde
General Requirements
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
Internal Delay Lines (SFU-SDL)
Translation stages (motorized)
Spec. Accuracy: <7µm

## Slide 10
SPICA-VIS Design Review, July 23-24, 2020
Spatial Filters Modules
General Requirements
SFU-INM Fiber Injection Optics
Optical system for the focalization of the beam at the fiber head level
SFU-OPF Optical Fiber
Optical Fiber for spatial filtering
SFU-RCM Fiber Reception Optics
Micro-Lens array for the beam collimation
Vgroove for beam set-up (3B–2B–4B–7B– B )
SCU Spectral Calibration Unit
Spectral sources for spectral Calibration
FBI Fiber Back Illumination
Light source to back illuminate to determine the reference axis of the fiber (on BCS)
Fiber Injection Optics (SFU-INM)
Lenses
Spec. Injection Level: >70% (CHARA beam)

## Slide 11
SPICA-VIS Design Review, July 23-24, 2020
Spatial Filters Modules
General Requirements
SFU-INM Fiber Injection Optics
Optical system for the focalization of the beam at the fiber head level
SFU-OPF Optical Fiber
Optical Fiber for spatial filtering
SFU-RCM Fiber Reception Optics
Micro-Lens array for the beam collimation
Vgroove for beam set-up (3B–2B–4B–7B– B )
SCU Spectral Calibration Unit
Spectral sources for spectral Calibration
FBI Fiber Back Illumination
Light source to back illuminate to determine the reference axis of the fiber (on BCS)
Optical Fibers (SFU-OPF)
Optical Fiber
Spec. Equalization in dispersion: 0.2mm

## Slide 12
SPICA-VIS Design Review, July 23-24, 2020
Spatial Filters Modules
General Requirements
SFU-INM Fiber Injection Optics
Optical system for the focalization of the beam at the fiber head level
SFU-OPF Optical Fiber
Optical Fiber for spatial filtering
SFU-RCM Fiber Reception Optics
Micro-Lens array for the beam collimation
Vgroove for beam set-up (3B–2B–4B–7B– B )
SCU Spectral Calibration Unit
Spectral sources for spectral Calibration
FBI Fiber Back Illumination
Light source to back illuminate to determine the reference axis of the fiber (on BCS)
Fiber Reception Optics (SFU-RCM)
Vgroove + Micro-Lens Array
Spec. Beam Spacing: 3B-2B-4B-7B-B

## Slide 13
SPICA-VIS Design Review, July 23-24, 2020
Spatial Filters Modules
General Requirements
SFU-INM Fiber Injection Optics
Optical system for the focalization of the beam at the fiber head level
SFU-OPF Optical Fiber
Optical Fiber for spatial filtering
SFU-RCM Fiber Reception Optics
Micro-Lens array for the beam collimation
Vgroove for beam set-up (3B–2B–4B–7B– B )
SCU Spectral Calibration Unit
Spectral sources for spectral Calibration
FBI Fiber Back Illumination
Light source to back illuminate to determine the reference axis of the fiber (on BCS)
Spectral Calibration Unit (SCU)
Spectral Lamp
Spec. Calibration lines spacing: <25nm between 2 lines

## Slide 14
SPICA-VIS Design Review, July 23-24, 2020
Spatial Filters Modules
General Requirements
SFU-INM Fiber Injection Optics
Optical system for the focalization of the beam at the fiber head level
SFU-OPF Optical Fiber
Optical Fiber for spatial filtering
SFU-RCM Fiber Reception Optics
Micro-Lens array for the beam collimation
Vgroove for beam set-up (3B–2B–4B–7B– B )
SCU Spectral Calibration Unit
Spectral sources for spectral Calibration
FBI Fiber Back Illumination
Light source to back illuminate to determine the reference axis of the fiber (on BCS)
Fiber Back Illumination (FBI)
White or laser source (motorized)
Spec. Simultaneous Illumination of the 6 fibers

## Slide 15
SPICA-VIS Design Review, July 23-24, 2020
Modules of the Spectrograph
Denis Mourard/Stéphane Lagarde
General Requirements
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
Interferometry/Photometry Splitter (IPS)
Beam splitter
Spec. Flux ratio Interferometry/Photometry: 90/10

## Slide 16
SPICA-VIS Design Review, July 23-24, 2020
Modules of the Spectrograph
Denis Mourard/Stéphane Lagarde
General Requirements
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
Photometry Transport System (PTS)
Lenses
Spec. Reimage the Vgroove output on the detector

## Slide 17
SPICA-VIS Design Review, July 23-24, 2020
Modules of the Spectrograph
Denis Mourard/Stéphane Lagarde
General Requirements
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
Interferometric Anamorphic System (IAS)
Cylindrical Optics
Spec. Anamorphic Ratio: >57

## Slide 18
SPICA-VIS Design Review, July 23-24, 2020
Modules of the Spectrograph
Denis Mourard/Stéphane Lagarde
General Requirements
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
Dispersion System (DIS)
Prism, Gratings (Motorized)
Spec. Spectral Resolutions: 140, 3000, 10000

## Slide 19
SPICA-VIS Design Review, July 23-24, 2020
Modules of the Spectrograph
Denis Mourard/Stéphane Lagarde
General Requirements
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
Camera Optics (CAO)
Lens
Spec. Sampling: ≥3pixels/fringe

## Slide 20
SPICA-VIS Design Review, July 23-24, 2020
Modules of the Spectrograph
Denis Mourard/Stéphane Lagarde
General Requirements
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
Detector (SDT)
ANDOR EMCCD IXON 888 (1024x1024)
Spec. DIT: ≤ 20ms
Frame rate ≥ 50 fps
Ok with 500 rows x 1024 columns

## Slide 21
SPICA-VIS Design Review, July 23-24, 2020
Global view
Denis Mourard/Stéphane Lagarde
General Requirements
VIS
STS Six Telescope Simulator
Non resolved sources for alignment cophasing and calibration
LDC-COM Longitudinal Dispersion Compensator
Chromatic OPD compensation between IR and Visible
CHARA
IR
SPICA-FT SPICA Fringe Tracker
Fringe tracking
LDC-VIS Longitudinal Dispersion Compensator
Chromatic OPD compensation for visible light
SPO-IMG Image Alignment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (image)
SPO-FOP SPICA Feeding Optics
Mirrors allowing the CHARA beams injection into SPICA
ADC Atmospheric Dispersion Compensator
Motorized double prisms for atmospheric refraction compensation.
BAS-PUP Pupil Adjustment Device
Motorized tip tilt mirrors for the CHARA/SPICA alignment (pupil)
PDC Polarization Delay Compensator
Motorized birefringent optics allowing the polarization phase delay compensation
BAS-TTT Tip Tilt Tracker
Fast tip tilt mirrors for the image tracking and fiber injection optimization
SHU Shutters
Individual motorized shutters to select beam(s)
BCS Beam Control System
Technical Camera
Control the CHARA/SPICA coalignment
Control the injection optimization in the fiber
SFU-SDL SPICA Delay Lines
OPD correction (between SPICA and the other CHARA instruments
SFU-INM Fiber Injection Optics
Optical system for the focalization of the beam at the fiber head level
SFU-OPF Optical Fiber
Optical Fiber for spatial filtering
SFU-RCM Fiber Reception Optics
Micro-Lens array for the beam collimation
Vgroove for beam set-up (3B–2B–4B–7B– B )
SCU Spectral Calibration Unit
Spectral sources for spectral Calibration
FBI Fiber Back Illumination
Light source to back illuminate to determine the reference axis of the fiber (on BCS)
IPS Inter/Photo Splitter
Beam Splitter allowing to split beams in 2 parts : Photometry and Interferometry
IAS Interferometric Anamorphic System
Cylindrical optics to optimize the signal sampling
DIS Dispersion System
Prism, Grating to disperse the signal
PTS Photometry Transport System
Transport the photometric channels to dispersive element
Image the pupil plane on the detector
CAO Camera Optics
Optical elements to combine the beams on the detector
SDT Detector
Signal detection
Photometry
Interferometry
SPICA-VIS
Alignment and Calibration Units
Beam Optimization Units
Spatial Filter Units
Spectro-Interferometer Units
Detector
Units external to SPICA-VIS